GLAST LAT ACD Failure Modes









(D1) RCH, June 11, 2001

Component
Function
Failure Type
or Cause
Failure Effect
Possible Mitigation
Performance after Mitigation
Remarks

ACD Subsystem
Background rejection
Elec. power loss
1) ACD unavailable to L1T; 2) Charged particle filtering much more difficult
=>Large reduction in LAT throughput
none
---
Category 2

ACD TEM
ACD interface
Total
"
Redundant TEM
Nominal
Category 2R

ACD Side
(1 of 4)
Background rejection
Elec. power loss
1) ACD less effective in L1T; 2) Charged particle filtering more difficult
=>Significant reduction in LAT throughput
none
---
Category 2 or3
(may not be possible)

Event board
HV generation and signal (9 or 18) processing
Elec. power loss
Efficiency loss for 9(18?) tiles (0.9997=>0.997)
Reduce thresholds for redundant PMT's
possibly some LAT efficiency loss
Category 3

Event board
HV generation and signal (9 or 18) processing
Command/data clock loss
1) 9(18?) VETO thresholds fixed; 2) 9(18?) CNO thresholds fixed; 3) Loss of PHA data for 9(18?) PMT; 4) Partial loss of VETO MAP
none
---
Category 3

Event board
HV generation and signal (9 or 18) processing
Command loss
1) 9(18?) VETO thresholds fixed; 2) 9(18?) CNO thresholds fixed; 3)Test mode unavailable
none
---
Category 3

Event board
HV generation and signal (9 or 18) processing
Loss of output data signal
1) Partial loss of VETO MAP; 2) Loss of PHA data for 9(18?) PMT's
none
---
Category 3

Event board
HV generation and signal (9 or 18) processing
Loss of TRIGACK
1) Partial loss of VETO MAP; 2) Loss of PHA data for 9(18?) PMT's
none
---
Category 3

HVPS failure
Activate 9 PMT's
Failed internal component
Efficiency loss for 9 tiles (0.9997=>0.997)
Switch to redundant HVPS(?)
Nominal(?)
Category 3or 4

ASIC 
PMT signal (9) processing
Elec. power loss
Efficiency loss for 9 tiles (0.9997=>0.997)
Reduce threshold for redundant PMT
Probably some loss of efficiency
Category 3 or 4

ASIC
PMT signal (9) processing
Channel loss due to internal failure
Efficiency loss for 1 tile (0.9997=>0.997)
Reduce threshold for redundant PMT
Probably some loss of efficiency
Category 4

PMT
Detect light from scintillator tile
Internal failure or loss of HV connection
Efficiency loss for 1 tile (0.9997=>0.997)
Efficiency loss for 1 tile (0.9997=>0.997)
Efficiency loss for 1 tile (0.9997=>0.997)
Category 4

PMT
Detect light from scintillator tile
Degradation (some expected)
Signal degradation
raise HV or lower thresholds
nominal
Category 4

VETO Signal
VETO function
Loss
Efficiency loss for 1 tile (0.9997=>0.997)
Efficiency loss for 1 tile (0.9997=>0.997)
Efficiency loss for 1 tile (0.9997=>0.997)
Category 4

Scintillator Tile Assembly
Detect charged particles via scintillation light
Light-exposure - penetration of light-tight wrap
Loss of functionality => some loss of DAQ filtering efficiency
none
---
Category 3 or 4

Fiber coupling
Conduct light from tile to PMT
Degradation due to vibration or aging
Signal degradation
raise HV or lower thresholds
Possibly some efficiency loss
Category 4

Fiber
Conduct light from tile to PMT
Break
Reduced efficiency in related portion of tile
raise HV or lower thresholds
Possibly some efficiency loss
Category 4

Analog sensor
Diagnostic information
Loss
Diagnostic information lost
none
---
Category 4









