A. Moiseev, Z. Arzumanian  

Updated 10/11/01      

                            Tile design test results

Task: study different designs of the ACD scintillating tile for the optimization. 

All scintillators tested are 1cm thick, readout by Hamamatsu R1635

1. Fiber spacing effect.

All tiles are BICRON, TYVEK wrapped, multiclad fibers

               Fiber spacing
             Light yield, relative units

   2 cm
15.9

   1 cm
18.1

    0.5 cm
22.0

   0.25 cm
20.2

   Continuous
23.3

2. Wrapping effect

Wrapping
Tile
Relative light yield

TYVEK
BICRON, 5mm fiber spacing, multiclad fibers
22.0

TETRATEC
 ----  “ “  ----
24.5

Poliester
 ----  “ “  ----
20.0





TYVEK
ElJen scintillator, 10mm fiber spacing, single clad fibers
15.3

TETRATEC
 ----  “ “  ----
17.2

Conclusion – TETRATEC gives ~ 10% of the light yield increase

3. Use of half of fibers (2 PMTs)

All tiles are BICRON, 5mm multiclad fiber spacing, wrapped in TETRATEC

Number of readout fibers
Relative light yield

All fibers 
24.5

Half fibers
11.9

Use of 2 PMTs reduces the light for each PMT by 50%

4. Aluminization of the fiber ends

Fiber ends
Tile
Relative light yield

Razor Cut
TYVEK, ElJen scintillator, 1cm single clad fiber spacing
14.1

Aluminized at GSFC
 ----  “ “  ----
15.3





Razor Cut
TETRATEC, BICRON, 5 mm multiclad fiber spacing
24.5

Mylar at the ends
---- “ “ ----
21.1





Razor cut
TYVEK, 1 cm multiclad fiber spacing
18.1

Aluminized in FNL
 ----  “ “  ----
21.2

 Conclusion – aluminization made at GSFC gives 5-7% of the light increase. Aluminization made in Fermilab, improves the light yield by ~ 17%

5. Fiber cladding

All tiles are ElJen scintillator, 1cm fiber spacing, TYVEK wrapped

Fibers
Relative light yield

Single clad
14.1

Multiclad
17.8

Conclusion:  cladding is important, gives ~ 25% of the light increase

6. Scintillator manufacturer

All tiles are 1cm multiclad fiber spacing, TYVEK wrapped

    Scintillator
Relative light yield

BICRON
18.1

ElJen1
17.8




Conclusion: Difference is within the measurement errors, scintillators perform similarly

7. Other different designs

Tile
Tested feature
Relative light yield

TETRATEC wrapped, BICRON
Light is transmitted to PMT by clear fibers viewing from the tile edge
5.5

BICRON, TYVEK wrapped
Light is collected by WSF glued to both tile edges without spacing
13.2



BICRON, TYVEC wrapped, 1cm multiclad fiber spacing
2 fibers in one groove
19.0 (to compare with 18.1 for one fiber)

BICRON, TYVEK wrapped, 1cm multiclad fiber spacing
Tile is made of 2  5mm thick tiles, fibers are glued between them in grooves cut in one tile
11.0

Design conclusion. It was found that the largest light yield will be provided by the following tile design:

a) 5 mm fiber spacing

b) TETRATEC as a wrapping material

c) Fiber ends are aluminized by the technology developed at Fermilab

d) Multicladding wave-shifting fibers are used

e) Scintillator manufacturer (Bicron or ElJen) does not matter much 

This design provides 50%-60% of the light increase compared with that obtained for BFEM tile design

