
Page 1 of 29 


	[image: image30.png]GARC Regitrs | GAFE Fogitrs | FunTest | AEM Regstes | Tem |

GAAE Regiter
SEREE e =

Reoites | PHA Theshad | ChamelErble | Mode/Stts

i Preset Ve 3] T Auorestavie wieal i Solntd
BEEEERD T~ Dispay Engineeing Ut e fosese
\ite Preset I~ selctél Read Al Read Selected
Cmd ommand Name: ata Value Wite [Read | W/R iror Status
St Cc d N: Data Valt Units: It It Status Enor Stat
| veropeLey e c
= T~ Paiy Eror
| hves f :I vos
~ ™ Crd Py Eror
| vsas I c I DataPaityEror
2 [ 2 - I CrdEror
se. o ﬂ MiPs »
W (Y ~| Cmd/Data ,—
I~ | HOLD_DEL&Y e = c in Enor
| VETO.WIDTH e = c
| Diagnostic Counters
T s c
I s c Vel Command
Courier
r [wmweeoeay o e » Diagnostic Sate
[ —
| MaPHa e = »
- [aoc_racn o 2 »

Cnd [ CrdName GAFE Chanmes e [Fesd | Wi
Select 171615141312 11108 8 7 55 4 3 210 I | I St
B [y | -
L [V Y= o o i C
N P Y S [l el i C





	LAT Document #
	Date Effective

	
	LAT-TD-01112-D1
	Draft 12/24/02

	
	Prepared by(s)
	Supersedes

	
	Dave Thompson
	None

	GLAST LAT Subsystem Technical Document
	ACD CM Document #
ACD-PLAN-000038
	

	
	Subsystem/Office

	
	Anticoincidence Detector Subsystem

	Document Title

	ACD Functional Test Plans (Comprehensive Performance Test)


Gamma-ray Large Area Space Telescope (GLAST)

Large Area Telescope (LAT)

ACD Functional Test Plans (Comprehensive Performance Test)
DRAFT 

CHANGE HISTORY LOG

	Revision
	Effective Date
	Description of Changes

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	


Prepared by:







Date:





Reviewed by:







Date:







 Jim La/Code 568

Approved by:







Date:







Ken Segal/Code 543

Approved by:







Date:







Glenn Unger/Code 561
Approved by:







Date:







Tom Johnson/Code 543

Approved by:







Date:







Alex Moiseev/Code 661

Approved by:







Date:







Bob Hartman/Code 661

Approved by:







Date:







Mike Amato/Code 556

Approved by:







Date:







Dave Thompson/Code 661

1. Purpose

The purpose of the Comprehensive Performance Test (CPT) is to define the most complete functional testing needed to verify the performance of the GLAST LAT ACD.  Other test plans, including Aliveness, Limited Functional, and Full Functional Tests, are subsets of this plan.  For each test plan there is a parallel detailed test procedure. 
2. Acronyms and Definitions 
	ACD
	The LAT Anti-Coincidence Detector Subsystem

	ADC
	Analog-to-Digital Converter

	AEM
	ACD Electronics Module

	ASIC
	Application Specific Integrated Circuits

	BEA
	Base Electronics Assembly

	CAL
	The LAT Calorimeter Subsystem

	CPT
	Comprehensive Performance Test

	DAQ
	Data Acquisition

	EGSE
	Electrical Ground Support Equipment

	EMC
	Electromagnetic Compatibility

	EMI
	Electromagnetic Interference

	ESD
	Electrostatic Discharge

	FM
	Flight Module

	FMEA
	Failure Mode Effect Analysis

	FREE
	Front End Electronics

	GAFE
	GLAST ACD Front End – Analog ASIC

	GARC
	GLAST ACD Readout Controller – Digital ASIC

	GEVS
	General Environmental Verification Specification

	GLAST
	Gamma-ray Large Area Space Telescope 

	HVBS
	High Voltage Bias Supply

	ICD
	Interface Control Document

	IDT
	Instrument Development Team

	I&T
	Integration and Test

	IRD
	Interface Requirements Document

	JSC
	Johnson Space Center

	LAT
	Large Area Telescope

	MGSE
	Mechanical Ground Support Equipment

	MLI
	Multi-Layer Insulation

	MPLS
	Multi-purpose Lift Sling

	PCB
	Printed Circuit Board

	PDR
	Preliminary Design Review

	PMT
	Photomultiplier Tube

	PVM
	Performance Verification Matrix

	QA
	Quality Assurance

	SCL
	Spacecraft Command Language

	SEL
	Single Event Latch-up

	SEU
	Single Event Upset

	SLAC
	Stanford Linear Accelorator Center

	TACK
	Trigger Acknowledge

	TDA
	Tile Detector Assembly

	T&DF
	Trigger and Data Flow Subsystem (LAT)

	TBD
	To Be Determined

	TBR
	To Be Resolved

	TSA
	Tile Shell Assembly

	TSS
	Thermal Synthesizer System

	TKR
	The LAT Tracker Subsystem

	VME
	Versa Module Eurocard

	WBS
	Work Breakdown Structure

	WOA
	Work Order Authorization


Comprehensive Performance Test (CPT) – a series of functional tests that exercises all functions of the ACD system and includes an absolute measurement of all the ACD tile efficiencies.

Full Functional Test (FFT) - a series of functional tests that exercises all functions of the ACD system except an absolute measurement of all the ACD tile efficiencies.

Limited Functional Test (LFT) - a series of functional tests that exercises all major functions  of the ACD system.

Aliveness Test (FFT) - a functional test that turns on the ACD in a nominal state and verifies basic operation of all channels. 

3. Applicable Documents

Documents relevant to the ACD Functional Test Plans include the following.

1. LAT-SS-00016, LAT ACD Subsystem Requirements – Level III Specification
2. LAT-SS-00352, LAT ACD Electronics Requirements – Level IV Specification
3. LAT-SS-00437, LAT ACD Mechanical Requirements – Level IV Specification
4. LAT-MD-00039-01, LAT Performance Assurance Implementation Plan (PAIP)
5. LAT-MD-00099-002, LAT EEE Parts Program Control Plan
6. LAT-SS-00107-1, LAT Mechanical Parts Plan
7. LAT-MD-00078-01, LAT System Safety Program Plan (SSPP)
8. ACD-QA-8001, ACD Quality Plan
9. LAT-TD-00720-D1 ACD Phototube Helium Sensitivity
10. LAT-DS-00740-1 Temperature Characteristics of ACD Photomultiplier Tubes
11. Response to RFQ 5-09742, Hamamatsu Photomultiplier Tube Proposal
12. LAT-MD-00404-01, LAT Contamination Control Plan

13. LAT ACD Subsystem Verification Table/Plan

4. ACD Functional Testing
Testing described in the Comprehensive Performance Test (CPT) includes operational tests that will be performed after the assembly of the ACD.  It assumes control of the ACD by an EGSE system with a high-level command system, details of which are TBD, and readout of ACD data through the EGSE data stream, with specific displays to be defined as needed.   

Because the CPT is a fairly exhaustive (i.e. time-consuming) test, it will be performed on a limited basis.  Subsections of the CPT will be used for subassembly testing, aliveness testing, limited functional testing, and full (but not comprehensive) functional testing. 

The CPT calls out specific tests that verify ACD Level IV requirements, as indicated in the ACD Subsystem Verification Table/Plan, and referenced by Level IV requirement number. 
Each test plan has a parallel test procedure that gives step-by-step instructions for the testing. Because each of the 194 channels is unique, much of the testing must be automated.
5. ACD Comprehensive Performance Test Plan
	Step
	Function
	Commands 
	Displays
	Analysis (reference to Level IV requirements verification in parentheses)

	Turn-On/Low Voltage Checkout

	1
	Power on primary AEM electronics 
	Turn on primary AEM power
	AEM Currents/voltages

	Compare with stored reference values


	2
	Power on FREE cards 
	Turn on low-voltage power
Send read command for all registers

194 x 11 GAFE registers,

12 x 42 GARC registers
	ACD Currents/voltages

24 FREE V

2 FREE I
Temperatures
Shell

BEA

PMT (4)

FREE (12)
	Compare with stored reference values (28 V, 3.3 V)
Verify default state as in ICD.
Verify on/off commands (5.14.3)

Record power in low-voltage state
Verify temperatures (5.9.10)

	3
	Configure nominal settings
	Set 194 x 6 GAFE registers.
Set 12 x 39 GARC registers.

Read back all.

NOT using zero-suppress mode.
	Command register

See Appendix C for sample display page for these. 
	Compare with stored reference configuration
Verify commanding (5.14.4, 5.14.5, 5.14.8)

Verify digital housekeeping (5.9.9)

	4
	Determine pedestals
	Send TACKs to each channel (194)
	PHA histogram
	Determine location and width of pedestal for each channel



	5
	Confirm test charge injection capability
	Test Charge injection – GAFE CONFIG,  TCI_DAC.
194 GAFEs x 2 levels
	Rates (AEM)

PHA histogram
	Verify test pulse injection (5.9.8) – VETO and both PHA ranges


	6
	Raise/lower VETO thresholds
	Change thresholds, VETO_DAC
Adjust charge injection, TCI_DAC

194 GAFEs 
	Command register

Rates
PHA histogram
	Rate v. commanded input
Compare with stored reference values
Verify test pulse injection level (5.9.8)

Verify adjustable VETO threshold (5.3, 5.14.4)

	7
	Raise/lower HLD threshold
	Change thresholds,
HLD_DAC

Adjust charge injection. TCI_DAC
194 GAFEs
	Command register

Rate
PHA histogram
	Rate v. commanded input
Compare with stored reference values
Verify adjustable HLD threshold (5.6, 5.14.5)



	8
	Compare VETO_AEM with VETO_Hitmap
	Generate TACK from ACD using charge injection. 
194 GAFEs
	Command register

Rates
	Rates v. threshold

Compare with stored reference values
Verify functioning of TACK (5.7, 5.14.7)

Compare rates (5.8.4, 5.9)



	9
	Verify low-range PHA
	Generate TACK from ACD using charge injection. 
Adjust charge to different levels, 
CONFIG_REG,

TCI_DAC
194 GAFEs
	Command register

PHA
	Verify functioning of TACK (5.7, 5.14.7)

Verify low-range PHA (5.9.3)

Calibrate PHA

	10
	Verify high-range PHA
	Generate TACK from ACD using charge injection 
Adjust charge to different levels

CONFIG_REG,

TCI_DAC

194 GAFEs
	Command register

PHA
	Verify functioning of TACK (5.7, 5.14.7)

Verify high-range PHA (5.9.3)

Calibrate PHA

	11
	Verify zero-suppress
	Generate TACK from ACD using charge injection 
Turn on zero-suppress.  
GARC mode.

PHA Thresholds (194 GAFEs),

12 GARCs

Adjust threshold. 

CONFIG_REG,

TCI_DAC

194 GAFEs
	Command register

PHA
	Verify functioning of TACK (5.14.7)

Verify PHA not transmitted for low values in zero-suppress mode (5.9.3)

	12
	Integral non-linearity
	Generate TACK from ACD using charge injection 
Adjust charge in steps. 

TCI_DAC

194 GAFEs x 64 steps
	Command register

PHA
	Verify PHA integral non-linearity (5.9.5)

	High Voltage Testing

	13
	Check light-tightness
	Turn on HVBSA, adjust to a low level, 
HVBS,

GARC mode,

then increase

HVBS.

12 GARCs x TBD steps.

Set VETO thresholds to nominal values

194 GAFEs
	Rates
HV monitor

PHA
	Compare rates with stored values.

	14
	Power on primary HVBS
	Turn on HVBSA, adjust to nominal values for tubes. 

HVBS,

GARC mode,

Set VETO thresholds to nominal values

VETO_DAC

194 GAFEs
	Rates

HV monitor
PHA
	Verify non-zero rates and PHA values (5.2).

Compare rates with stored values.

	15
	Raise/lower HV
	Change HV
HVBS

12 GARCs
	Rates
Command register
	Rates v. HV

Compare with stored reference values
Verify HV commanding (5.14.2)

	16
	Collect PHA spectra (short)

Self-triggered gain test 
	Generate TACK from ACD
194 GAFEs
	PHA spectra
	Calculate pedestals

Fit spectra with expected shape

Compare with stored reference values
Verify pulse digitization (5.9.3)

See Appendix A.

	17
	Determine absolute thresholds
	Generate TACK from ACD
Adjust VETO thresholds

VETO_DAC

194 GAFEs
	PHA spectra
	Compare with stored reference values
Compute thresholds in MIPs (5.3)

	18
	Re-check pedestals
	Send TACKs to each channel (194)
	PHA histogram
	Determine location and width of pedestal for each channel



	19
	Measure false VETO
	Lower HV

HVBS
	Rates
	Verify false VETO rates (5.4)

	20
	Test discriminator masking
	Change discriminator mask
PHA EN0

PHA EN1

12 GARCs

194 GAFEs

2 levels
	Rates
	Compare with stored reference values
Verify discriminator masking (5.8.7)

	21
	Limit and state change checking

	Enable limit checking and state change checking by EGSE
	Messages for state changes, highlighted parameter values if out of limit 
	Compare currents, voltages, temperatures with a limit table. Compare register states with a nominal configuration table. 

	22
	Gain calibration test

Collect PHA spectra (long)


	Generate TACK from ACD
	Gated PHA spectra
	Trigger on VETO signal from any tile, use this to gate PHAs for all other tiles in a selected pattern

Calculate pedestals

Fit spectra with expected shape

Compare with stored reference values
See Appendix B

	23
	SAA Mode test
	SAA signal
SAA
	Rates, HV monitor
	Compare with stored reference values
Verify command works , mode change takes place, and notification done  (5.14.9, 5.14.10)

Measure power in this configuration. 

	24
	Power measurement
	N.A.
	Currents, voltages
	Measure total power when operating  (5.16)

	Repeat HV tests with secondary HV Bias supplies

	25
	Check light-tightness
	Turn on HVBSB, adjust to a low level, 

HVBS,

GARC mode,

then increase

HVBS.

12 GARCs x TBD steps.

Set VETO thresholds to nominal values

194 GAFEs
	Rates

HV monitor

PHA
	Compare rates with stored values.

	26
	Power on secondary HVBS
	Turn on HVBSB, adjust to nominal values for tubes. 

HVBS,

GARC mode,

Set VETO thresholds to nominal values

VETO_DAC

194 GAFEs
	Rates

HV monitor
PHA
	Verify non-zero rates and PHA values (5.2).

Compare rates with stored values.

	27
	Raise/lower HV
	Change HV
HVBS

12 GARCs
	Rates
Command register
	Rates v. HV

Compare with stored reference values
Verify HV commanding (5.14.2)

	28
	Collect PHA spectra (short)

Self-triggered gain test 
	Generate TACK from ACD
194 GAFEs
	PHA spectra
	Calculate pedestals

Fit spectra with expected shape

Compare with stored reference values
Verify pulse digitization (5.9.3)

See Appendix A.

	29
	Determine absolute thresholds
	Generate TACK from ACD
Adjust VETO thresholds

VETO_DAC

194 GAFEs
	PHA spectra
	Compare with stored reference values
Compute thresholds in MIPs (5.3)

	30
	Measure false VETO
	Lower HV

HVBS
	Rates
	Verify false VETO rates (5.4)

	31
	Test discriminator masking
	Change discriminator mask
PHA EN0

PHA EN1

12 GARCs

194 GAFEs
	Rates
	Compare with stored reference values
Verify discriminator masking (5.8.7)

	32
	Gain calibration test

Collect PHA spectra (long)


	Generate TACK from ACD
	Gated PHA spectra
	Trigger on VETO signal from any tile, use this to gate PHAs for all other tiles in a selected pattern

Calculate pedestals

Fit spectra with expected shape

Compare with stored reference values
See Appendix B

	33
	SAA Mode test
	SAA signal
SAA
	Rates, HV monitor
	Compare with stored reference values
Verify command works , mode change takes place, and notification done  (5.14.9, 5.14.10)

	34
	Power measurement
	N.A.
	Currents, voltages
	Measure total power when operating  (5.16)

Compare with primary HVBS power. 

	Repeat some tests with secondary AEM 

	35
	Power on secondary AEM electronics 
	Turn on secondary AEM power
	AEM Currents/voltages

	Compare with stored reference values


	36
	Power on FREE cards 
	Turn on low-voltage power
Send read command for all registers

194 x 11 GAFE registers,

12 x 42 GARC registers
	ACD Currents/voltages

24 FREE V

2 FREE I

Temperatures
Shell

BEA

PMT (4)

FREE (12)
	Compare with stored reference values (28 V, 3.3 V)

Verify default state as in ICD.

Verify on/off commands (5.14.3)

Verify temperatures (5.9.10)

	37
	Configure nominal settings
	Set 194 x 6 GAFE registers.

Set 12 x 39 GARC registers.

Read back all.
	Command register

See Appendix C for sample display page for these. 
	Compare with stored reference configuration
Verify commanding (5.14.4, 5.14.5, 5.14.8)

Verify digital housekeeping (5.9.9)

	38
	Confirm test charge injection capability
	Test Charge injection – GAFE CONFIG,  TCI_DAC.

194 GAFEs x 2 levels
	Rates (AEM)

PHA histogram
	Verify test pulse injection (5.9.8) – VETO and both PHA ranges



	39
	Check light-tightness
	Turn on HVBSA, adjust to a low level, 

HVBS,

GARC mode,

then increase

HVBS.

12 GARCs x TBD steps.

Set VETO thresholds to nominal values

194 GAFEs
	Rates

HV monitor

PHA
	Compare rates with stored values.

	40
	Power on primary HVBS
	Turn on HVBSA, adjust to nominal values for tubes. 

HVBS,

GARC mode,

Set VETO thresholds to nominal values

VETO_DAC

194 GAFEs
	Rates

HV monitor
PHA
	Verify non-zero rates and PHA values (5.2).

Compare rates with stored values.

	41
	Raise/lower HV
	Change HV
HVBS

12 GARCs
	Rates
Command register
	Rates v. HV

Compare with stored reference values
Verify HV commanding (5.14.2)

	42
	Collect PHA spectra (short)

Self-triggered gain test 
	Generate TACK from ACD
194 GAFEs
	PHA spectra
	Calculate pedestals

Fit spectra with expected shape

Compare with stored reference values
Verify pulse digitization (5.9.3)

See Appendix A.

	
	
	
	
	

	

	Efficiency Testing 

	43
	Efficiency test – requires ACD in three different orientations. (very long runs)
	Generate TACK from ACD
	Gated PHA spectra
	Trigger on VETO signal from any tile, use this to gate PHAs for all other tiles in a selected pattern

Calculate pedestals

Fit spectra with expected shape

Compare with stored reference values that have absolute efficiency calibration.  (5.2)

See Appendix B

	44
	Determine latency
	External gate from muon telescope to generate TACK

Delay box
	Rates, PHA spectra
	Add and subtract delay until VETO signals and PHA signals are no longer seen (5.8.1)

	45
	High-rate protection
	Configure AEM for limit checking on tile rates
	Rates, HV
	Use radioactive source to force tile rate above 10 KHz, watch for SAA command to reduce HV.

	
	
	
	
	


6. Redundancy (Cross-Strapping) Test Plan
Although the ACD is not fully redundant, there are parts of the system that have independent parallel components, namely the AEM and the HVBS (note that all FREE cards and all phototubes are used for normal operation).  Tests should involve these independent subsystems under these guidelines:

· We wish to exercise the primary and secondary systems approximately equally, in order to gain confidence that we understand any possible performance differences and to assure that “infant mortality” of the systems is not an issue.

· One approach to achieving this goal:  A functional test should not always start with the primary AEM and the primary HVBS, but cycle through the four combinations (Primary-Primary, Primary-Secondary, Secondary-Primary, Secondary-Secondary).  Separate procedures may be needed. 

· If both primary and secondary AEMs are being tested within one procedure, the alternate (whichever it is) does not need to repeat all the tests done on the first AEM.  The VETO signals are distinct, as are the thermistors, but commands and serial data (?) go through the same circuitry on the GARC.  For this reason, a sample of commands and data will be adequate to confirm performance. 
7. Shorter Versions of Functional Testing

Aliveness test – steps 1, 2, 3, 13, 14, 16.

Abbreviated functional test – steps 1, 2, 3, 4, 5, 8, 9, 11, 13, 14, 16, 17, 21, 22
Full functional test – all steps except 43, 44 and 45. 

Appendix A - Self-Triggered Mode Information (Steps 16 and 28 of CPT)

The Self-Triggered Mode of testing flight ACD. One way to do the gain calibration test is to look at all tile histograms in muon run self-triggering mode, meaning that all signals appeared in given tile will be used for the analysis. The advantage of this approach is that for ~ 1 hour of instrument running there will be from 3,000 to 15,000 events in the histogram (depending on the tile) which provides a very reliable and precise peak position determination. The disadvantage of this approach is that the histograms for some tiles, especially for the side tiles, will be very dependent on the muon flux angular distribution. This is because no external triggering will be used, so the angular range of particles causing the triggering will be 2( for every tile, which in the convolution with the incident flux angular dependence could cause the uncertainty if the latter one varies. Simulated histograms for self-triggering mode, are presented in fig. 1 and 2. 
[image: image1.jpg]


[image: image4.jpg]tiles.

5x 320

calorimeter





                Tile T11 – middle tile in the edge row     Tile T12 – next to the central tile

       Fig. 1 – Top tiles in self-triggering mode

[image: image5.jpg]3

0

1

10

Tile T11 triggered by Z23

@215 1155

T T T T T TH ]

et

oo o518]
605.9

205.3

LINEANCC|
519
9846
o765




[image: image6.jpg]1

10

Tile T12 triggered by X41+X42+Z41+242

Q15 1158

£

1000

it

592
H——a
21£%]

9 ¥ T

6064
oo 75391
1373

100

1400

1600




          Tile T13 – top center tile                                    Tile T17 – diagonal from the central tile

[image: image7.jpg]Tile 13 triggered by x41+x42+z41+4242 020015 1134

O R A e ) AT SRR SRR RS ER T
Earries 524

Wlean
RMS

X/ndf.
BL
"

B




[image: image8.jpg]Tile t17 triggered by X41+X42+Z41+242

Q015 12,08

30

£

1000

BT
Egries--
Wlean
RMS

1619

;2]
B

100

1400

1600



[image: image9.jpg]0/02/15 1246
tile Z11 triggered by T2+T3+T7+T8+T12+T1 3+opposite side

T T T T BT T IRETH
Extries 127
e T
RMS 270.6

X/ ndf 2785 ¥ 42
r L R N [0
;2] 12.79

B3 2303

1000 D200 10 00




Tile Z11 – end tile in the  top row       Tile Z13 – middle tile in the          Tile Z21 – end tile in the

                                                              top  row                                           second row

   Fig. 2 – Side tiles

[image: image10.jpg]10

Q15 1221
ig. by T7+T8+T12+T13+T17+T18+opposite side

REERRA i Ea

T T a7

B

Harries 149
L e S e
RMS 2806
X/ ndf 3547 ¥ 46
Bl 5299
B o4

145.5

w0l Do 100

1600



[image: image11.jpg]Q15 1314
Tile Z21 trig. by T2+T3+T7+78

SERTAELRAEECAEERI RS TARAIL AEERERTAREAI LR T
Entries 214
Nlear

RMS, |

X/ ndf

BL

L
B




[image: image12.jpg]s

1

10

75

Tile 222 trig. by Ta+T7+T8+T13 %0 B

T T T T T TH T T
[ — H T RS T !
Mean 696.5

RMS_ | 2697
X/ ndf 3748 ¥ 34
Bl 5664

B2 i59.63]
B 123.0

a0 100 D0 Lo 1600 1200




[image: image13.jpg]i

10

Tile 231 trig. by T2+T34T74T84T124T13 =% B4

TR Lo e [ epnt [ 6 T oI o
Brries S
ol i L i w230

RMS 3417
I — ndfl 3727 36
L T4l

L ueT3
i — 11883

a0 100 D0 L0 300 1300




How sensitive is the self-triggering mode to the gain change? The gain change by 5% was simulated. Similar thing was experimentally tested and proven with muons on BFEM. The simulation run identical to that given above, but with the gain of 0.95 was performed. The MIP peak positions for these runs are given in the Table 3 along with the statistics in each corresponding histogram. 

	Tile
	Statistics for 40 min
	Peak position for gain = 1
	Peak position for gain = 0.95
	Ratio

	T11
	10,650
	498.9
	472.7
	0.947

	T12
	10,680
	502.2
	476.4
	0.949

	T13
	10,828
	497.7
	471.6
	0.947

	T17
	10,651
	497.4
	471.6
	0.948

	Z11
	3,631
	785.6
	748.9
	0.953

	Z13
	3,601
	793.0
	756.7
	0.954

	Z21
	2,547
	784.6
	748.8
	0.954

	Z22
	2,512
	776.2
	729.7
	0.940

	Z31
	1,930
	739.8
	711.2
	0.961

	Z32
	1,944
	765.5
	735.8
	0.961

	Z41
	9,658
	776.8
	738.0
	0.950


Obtained results demonstrate that the approach is quite sensitive on the required level of sensitivity (5% of the gain relative change)

Appendix B – Information about Full ACD Gain Calibration Testing (Steps 22, 32, and 43 of CPT)
[image: image14.jpg]s

1

ns

10

75

@5 1535
Tile Z32 trig. by T2+T3+T7+T8+T12+T13

BT

IRERAERTE
302

Entries

Mean 8512
RMS T T 3428
X/ndf 5407 4L
BL 2 & 73611
;2] 156.72

175.0




[image: image15.jpg]QA5 15.45

Tile Z41 trig. by T3+T8+T134T18+723
T T T T T

3 [

PREEcaE RIS IiEEA R KR RIERTE]
Entries 551
Mean 750.1
RMS T T 270
X/ndf 6775 ¥ 40
BL & i846:27
;2] 154.9

0

B3

126.8

1

10




[image: image16.jpg]@03/06 1653

996 events

20




Test of the Flight ACD .  The task is to find for each particular tile what tiles can be used to trigger it.  The trajectories should be as normal to the tested tile surface as possible, with reasonable statistics to be collected from the cosmic muons.  The tile numbering used in the simulations is shown in Fig. 3.  The simulation run corresponded to approximately 40 minutes of ACD running time, with 10-11 thousand triggers collected for each of top tiles.  For each tile, the triggering tiles were carefully selected, and corresponding histograms are shown in figures below.  For reliable fitting and MIP peak position determination, approximately 1,500 events are desirable in the histogram.  Looking at the histograms, we see that the most difficult tiles to calibrate will be the upper side tiles (fig. 5), which will require 6-7 hours to get ~1,500 events.  Limited calibration can be done within ~4 hours Steps 19 and 29). 

Full calibration (Step 32) requires higher statistics with most particles normal to the tiles being tested.  The approach is to change the orientation of the ACD:  (1) top face horizontal; (2) one pair of sides horizontal; (3) opposite pair of sides horizontal.  

[image: image17.jpg]511 events 02103006 1649

0

o 2l




[image: image18.jpg]Tile 5 BFEM, 374 events

T T B
Entries
Meah
RMS§

/0
Pl

P2

P3





               Tile T11 – middle of top edge                                       Tile T12 – next to the top central tile

[image: image19.jpg]Tile T11, self—trig, gain=1.

D] mar
Entsies. 1p799-]
Mesah 6244
RMS 458
X/ ndf 4330 7 95
PIY T TTTAg8
P2 {883

P3 109.7

1000 D0 10 k0o



[image: image20.jpg]Tile T12, self—trig, gain=1.

D]
Entries

Mesh
RMS

X/ ndf
Pl

P2

P3

w0 100 D0 100

1600





                       Tile T13 – top center tile                                Tile T17 – diagonal from the top central tile

Fig.4 - Top tiles

Fig. 5 - Side tiles

[image: image21.jpg]e

30

100

Tile T13, self—trig, gain=1.

RERJEREAEERA B

EfithgS
Mesh
RMS

ar

X/ ndf
PL

P2
P3




[image: image22.jpg]Tile T17, self—trig, gain=1.

D]
Entries

Mean
RMS§

X/ ndf
Pl

P2

P3





    Tile Z11 – end tile in the side top row                                Tile Z13 – middle tile in the side top row

[image: image23.jpg]0

0

50

0

10

Tile Z11, gain = 1, self—triggered

Q0307 14.18

B

T a7

“Brnedes T T 7363t
Mean 98l.4
BMS 4384
R ndft F60:8- i 97—
BL {785.6
;2] 192.4

12987




[image: image24.jpg]0

0

50

0

10

Tile Z13, self—trig, gain=1.

D]
Eiitries
Mesan
RMS

Imar
80T
972.3
437.3

X/ ndf
Pl

PRt

2063 /£ 97
793.0
198:5+
8L8





         Tile Z21 – end tile in side second row                  Tile Z22 – second tile from end in side second row

[image: image25.jpg]5

3

0

1

10

Tile Z21, gain = 1, self—triggered

Q0307 14.20

1000

F|

Entries 2548
Mean 974.0
RMS T T At
X/ ndf

BL 2

100 100

1600

1202




[image: image26.jpg]Tile Z22, self—trig, gain=1.

D] ! mar
) - i Entsies..... 2512-|
Meah 963.3
RMS 456.8
X/ndf 1560 / 97
PLTTTTTTT TR

258

320.8

0

10

a0 100 100 100 00 1a0 200




         Tile 31 – end tile in side third row                         Tile Z32 – second tile from end in side third row

[image: image27.jpg]Tile Z31, gain = 1, self—triggered Q37 1421

RN EaRABASN RARARAGN B i NEAN EASARERIL

Harries o0

3

0

1

10

Mean 9415
NS~ ST
&/ndf 3
Bl 739.8
" 95.03

B3

298.8





          Tile Z41 - Long tile, side bottom row

How many events do we need in the histogram for given peak position uncertainty? The common mathematical approach to this estimate is complicated by the high variability of the particle paths in the tile and desirable lowering the number of events needed. The simulations seem to the appropriate way to do this analysis. I did the following – using sea level muon flux in the simulations, the simulated pulse height distribution was fitted by Landau distribution to find the peak position. This was repeated for 10 sets of approximately 2,500, 1,000, and 500 events in the histogram and the mean value and standard deviation (() was determined for each of these sets of 10 runs. This was done for a tile on the top of ACD (where most of the muons hit the tile on or around normal incidence). The results are given in Table 2. The examples of the pulse height distribution for the top tile, 996 events and 511 events in the histogram are given in fig. 8 and fig.9 (column 3 and 4 in the table) respectively. The histogram for the side (“bad”) tile and 374 events is shown in fig. 10 (5-th column in the table). It is seen that the precision of the peak position fitting is surprisingly high, even for such a small statistics as ~ 500 events are and for the “bad” tile (the side one) there are a large variation of the incident muon angles, and consequently the muon paths in the tile. We are looking for the change in the light yield on a level of 5% and more.  For the efficiency test, we will want higher statistics with more particles moving normal to the tiles. 
Table 2   Simulations of the peak position determination precision

	Fitted MIP peak position
	Top tile, ~2,500 events
	Top tile, ~1,000 events
	Top tile, ~500 events
	Side tile, ~400 events, gain=1
	Side tile, ~400 events, gain=0.95

	
	330.8
	329.9
	332.2
	487.3
	467.4

	
	333.0
	326.8
	338.2
	495.2
	467.2

	
	336.7
	336.6
	324.5
	510.3
	462.8

	
	340.9
	338.4
	331.5
	481.3
	489.9

	
	330.7
	335.0
	340.9
	482.1
	478.9

	
	334.9
	330.9
	348.4
	491.9
	539.5 (

	
	331.0
	336.0
	326.8
	509.5
	447.1

	
	336.8
	339.1
	343.0
	510.9
	482.1

	
	331.2
	339.5
	336.5
	511.7
	482.5

	
	336.2
	334.8
	350.8
	521.1
	476.4

	Mean ((
	334.2(3.4 (1%)
	334.7(4.3 (1.3%)
	337.3(8.7 (2.6%)
	500.1(13.8 (2.8%)
	478.7(24 (5%)


[image: image28.jpg]160

10

10

100

0

0

Tile Z41, self—trig, gain=1.

T T Il
Entries
Mesan
RMS:

X/ ndf
Pl

P2

1000 100

1400 1600



[image: image29.jpg]JajawpIoED

z w51
- B
sz ez 34 wz Tz
sz ez [:23 wz ez | [
Ty
stz 1z
L STL
@y oy w1
1 L L
21 B 11 61
s o b oL

matExs

sm





[image: image2.wmf]
Fig. 8   Top tile, 996 events       Fig.9  Top tile, 511 events     Fig. 10  side tile, 374 events



Appendix C – Sample Display Page for Command Registers

[image: image3]
Appendix D – Calibration Information to be Derived from Functional Testing

Table 1 – Calibration Parameters for Low-Energy Range

	Tile/Ribbon
	Tube
	Rate (Hz)
	Pedestal
	Elec. Resp.
	MIP Pos.
	MIP FWHM
	Thresh. command
	Thresh. (derived)
	HV command
	 HV

readout
	Effic. (derived)

	000
	0
	350
	125
	1000
	1200
	300
	440
	450
	1050
	1050
	.9998

	
	1
	 
	 
	 
	 
	 
	 
	 
	 
	 
	

	001
	0
	 
	 
	 
	 
	 
	 
	 
	 
	 
	 

	
	1
	 
	 
	 
	 
	 
	 
	 
	 
	 
	


.

.

.

	602
	0
	350
	125
	1000
	1200
	300
	440
	450
	1050
	1050
	.9998

	
	1
	 
	 
	 
	 
	 
	 
	 
	 
	 
	

	603
	0
	 
	 
	 
	 
	 
	 
	 
	 
	 
	 

	
	1
	 
	 
	 
	 
	 
	 
	 
	 
	 
	


Note:  Electronic Response for the low-energy range is defined as the channel in which a 0.64 pC charge appears.

Table 2 – Calibration Parameters for High-Energy Range

	Tile/Ribbon
	Tube
	Rate (Hz)
	Pedestal
	Elec. Resp.
	CNO Pos.
	CNO FWHM
	Thresh. command
	Thresh. (derived)
	HV command
	 HV

readout
	

	000
	0
	350
	125
	1000
	1200
	300
	440
	450
	1050
	1050
	

	
	1
	 
	 
	 
	 
	 
	 
	 
	 
	 
	

	001
	0
	 
	 
	 
	 
	 
	 
	 
	 
	 
	 

	
	1
	 
	 
	 
	 
	 
	 
	 
	 
	 
	


.

.

.

	602
	0
	350
	125
	1000
	1200
	300
	440
	450
	1050
	1050
	

	
	1
	 
	 
	 
	 
	 
	 
	 
	 
	 
	

	603
	0
	 
	 
	 
	 
	 
	 
	 
	 
	 
	 

	
	1
	 
	 
	 
	 
	 
	 
	 
	 
	 
	


Note:  Electronic Response for the high-energy range is defined as the channel in which a 31 pC charge appears.

Table 3 – Linearity Response for Low-Energy Range
	Tile/
Ribbon
	Tube
	Step 0
	Step 1
	Step 2
	Step 3
	• • •
	Step 61
	Step 62
	Step 63

	000
	0
	 
	 
	 
	 
	 
	 
	 
	 

	
	1
	 
	 
	 
	 
	 
	 
	 
	 

	001
	0
	 
	 
	 
	 
	 
	 
	 
	 


.

.

.
	603
	0
	 
	 
	 
	 
	 
	 
	 
	 

	
	1
	 
	 
	 
	 
	 
	 
	 
	 


Table 4 – Linearity Response for High-Energy Range

	Tile/

Ribbon
	Tube
	Step 0
	Step 1
	Step 2
	Step 3
	• • •
	Step 61
	Step 62
	Step 63

	000
	0
	 
	 
	 
	 
	 
	 
	 
	 

	
	1
	 
	 
	 
	 
	 
	 
	 
	 

	001
	0
	 
	 
	 
	 
	 
	 
	 
	 


.

.

.
	603
	0
	 
	 
	 
	 
	 
	 
	 
	 

	
	1
	 
	 
	 
	 
	 
	 
	 
	 











Fig. 11  For the full efficiency test, the ACD will need to be rotated as shown and into a similar configuration with the orthogonal sides horizontal.





Fig. 3 – Tile numbering used for these simulations.  The standard tile numbering convention will be used for actual data and testing. 





















































     Tile Z22 – second tile from end in        Tile Z31 – end tile in 3-rd row   Tile Z41 – Long tile 


       side second row                                                                                        bottom row














Fig. 5








Draft

_1077024992.doc

