Directive No.
GLAST-ACD-B Test

Approved By: Tom Johnson    
Effective Date: November 22, 2002


Title:  Instrument Manager
Expiration Date: November 15, 2005


Authors: Bill Reaves, Tom Riley

[image: image6.png]



Responsible Office: GSFC / Code 661 / GLAST ACD
Title:  TDA-PMT-Resistor Network End-to-End Thermal-Vacuum Test Procedure

1. Purpose

This Document defines the Engineering Tests to be performed on two Tile Detector Assemblies (TDAs) and Photo Multiplier Tube (PMT)/Resistor Network Assemblies in a thermal vacuum chamber for the Gamma-ray Large Area Space Telescope (GLAST) AntiCoincidence Detector (ACD).

2. Scope

This document describes the procedures to verify the design of the TDA, Clear Fiber, and PMT assembly) under thermal-vacuum conditions, including high temperature operational & survival, low temperature operational & survival, and thermal cycling.

3. Acronyms

	ACD – AntiCoincidence Detector

	GLAST – Gamma-ray Large Area Space Telescope

	TDA – Tile Detector Assembly

	T-Vac – Thermal Vacuum

	TTA – Triggering Tile Assembly

	PMT Photo multiplier Tube

	RN- Resistor Network


4. Applicable Documents

This test requires that the following documents completed.

	Tile Detector Assembly (TDA) TEST B PROCEDURE

	Tile Detector Assembly (TDA) TEST A PROCEDURE

	Tile Detector Assembly (TDA) Thermal Vacuum Test Report

	PMT/Resistor Network Test Procedure

	


5. Safety Precautions and Warning Notes 

5.1. The PMT/RN assemblies operate at approximately 1000V DC.  Do Not operate any PMT/Rn at a pressure less than 1 atm. or greater than 1*10 –5 torr, nor during transition from 1 atm. to vacuum (<1*10 –5 torr).

5.2 All personnel who come in direct contact with the hardware under test shall wear charge conducting wrist straps. Each strap shall be connected to facility ground. 

6. References

TDA Test B Procedure

      PMT/Resistor Network Test Procedure

7. Resources

7.1. Tools

	Tools
	Make, Model, Vendor

	TBS
	TBS


7.2. Materials

	Materials
	Make, Model, Vendor

	Cleaning materials
	TBS


7.3. Equipment

	Equipment
	Make, Model, Vendor, S/N

	High Voltage Power Supply
	Tennelec TC 952

	Digital Multi-Meter
	Fluke

	High Voltage Probe
	Fluke 80K-40 HV Probe

	Thermal Vacuum Chamber
	Building 4

	TDA Signal Pre Processor Rack
	LeCroy

	VME Box
	Data Design Corp., SCICrate

	PC, monitor, keyboard, printer
	TBS

	High Resistance Tester
	Hewlett Packard HP 4329A

	TDAs with Clear Fiber
	Fermi Labs  S/N 01,02, 03

	Low Voltage Power Supply
	Topwood 6302A

	Oscilloscope
	Tectronic TDS754C


7.4. Fixtures

	Fixtures
	Drawing #

	Support for TDA under test
	Figure 2

	Support for TTA Tiles
	Figure 2

	Test Cables
	TestbCab.doc


8. Personnel

The test will require one test conductor between 7:00 am to 5:00 pm, and two test conductors between 5:00pm and 7:00 am.

All test conductors will have training in the use of the ACD TDA test equipment and the operation of the Building 4 thermal-vacuum chamber. 

The Thermal Engineering Branch will provide personnel to perform specific operations related to the thermal vacuum chamber and on-site equipment in building 4.

9. Location

This Thermal Vacuum Test will be conducted the Thermal Engineering Branch’s 36-inch vacuum chamber located in Building 4. Room 191 at GSFC.

10. Test Procedure

10.1. Set Up Test Equipment

[image: image1.wmf]Tom Riley

6-0712

11/22/2002

Btest1c.skf

HVPS

Cosmic Rays

Res 

& PMT

2 Flight Tiles

Meter

Pre Processor Box

Stanford Data Logger

PC

Monitor

(set up only)

Vacuum Chamber

Figure 1 -- Engineering "B" Test Configuration

HV Probe

Printer

Word

Excel

CD Burner

3.5 Disk

c

6 - Single Conductor HV Feed-Thru

Scope

Cast Block

600 -

 1250 V

Non-Flight Rail

t

LVPS

t

Thermisters

Keyboard

Hi-Res

Tester

tp1 Rtn

1

(Flight Mount,

2 PMT's each)

Base

M1

M2

Cable 3

Cable 1

Cable 2

Cable 4

Gnd

Gnd

Rtn

4

5

1 Non-Flight

 Tile

Trigger

VME Box

Safety

Cover

Preamps

(LeCroy)

Cable 6

Cable 7

2

3

1

4

5

2

3

t

t

t

t


10.1.1. Clean Test Area

Use shop vacuum cleaner on floor and household cleaner on bench tops, and Windex on computer screens.


Figure 2

10.1.2. Set Up Internal Unit Under Test

Using the TDA Test B Procedure. Set-up test equipment in and around T-Vac chamber as shown in Figure 1.  

10.1.2.1. Position Unit Under Test 

Place the units under test on the TDA test fixture in the vacuum chamber.  Attach PMT/Resistor network assemblies and fixture on the thermal vacuum chamber baseplate.  

10.1.2.2. Make Internal Connectors

Attach safety ground wires to rail and stand.  Make internal electrical connectors, internal high voltage, and signal cables, as shown in Figure 1 and Appendix A.

Enter the serial number of the PMT assembly, the number of the internal High Voltage, and the high voltage feed-thru as a clock position (as seen from outside) in the following table.
The High Voltage common must the on the center feed-thru.

	PMT Serial Number
	Internal HV Cable Number
	Feed-thru Clock Position

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	


Enter the serial number of the PMT Resistor assembly, the internal cable number and the external cable number, and the Preamp Input number in the following table.  The internal cables may be marked with bands of kapton tape.  

	PMT Serial Number
	Internal Cable
	External Cable 
	Preamp Input

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	


10.1.2.3. Site thermocouples

Attach thermocouples to test hardware at locations designated  below, and as shown if Figure 2.  

Thermocouple  #





Location

1


   



PMT Assembly

2


   



PMT
1

            3


   



PMT
2

4


   



PMT
3

5


   



TDT PMT


6


   



TDA 01 Connector

7


   



TDA 02 Connector

8


   



Panel (Top)

9


   



TDA Tile Bottom



10






TDA Tile Top

10.1.3. Set Up External Test Equipment

Using the TDA Test B Procedure Section 3., set up external test equipment outside the vacuum chamber as shown in Figure 1.

10.1.3.1. Aliveness Test with Scope

Turn On Power supplies and set to initial voltage using the calibrated digital multimeter.  Using the TDA Test Procedure Section 3.2 thru Section 3.17 use the oscilloscope to demonstrate that each PMT assembly produces data events.

10.1.3.2 Disconnect The High Voltage 

Turn off high voltage power supplies and de-mate the high voltage connections at the high voltage supplies.

10.1.3.3    High Resistance Test

Using the PMT/Resistor Network Test Procedure, measure the resistance between the high voltage 

+  input and high voltage return (shield for high voltage connectors), for each PMT/ Resistor Network.  Log resistances below.

Warning:  Failure to connect the High Voltage connections properly could damage flight equipment.

PMT S/N

(+)Probe
(-) Probe
Expected

Resistance

_____

Center Conductor
Shield

886 +/-88Mohms
_________ Mohms 

_____  
Center Conductor
Shield

886 +/-88Mohms
_________ Mohms

_____

Center Conductor
Shield

886 +/-88Mohms
_________ Mohms

_____

Center Conductor
Shield

886 +/-88ohms
_________ Mohms

10.1.3.4      Reconnect HV Cable

Reconnect the External High Voltage and check that the common return is on the center feed-thru.

Warning:  An improper connection of the External High Voltage Cable could damage Test Hardware.

10.1.3.4. Close Vacuum Chamber  

This step is to be performed by Code 545 personnel.

10.1.3.5. Functional Test B

Run TDA Test B Procedure (Section 4. ) with the chamber closed to demonstrate that all cables are connected properly and all PMT are operating.

10.2. Vacuum Testing

Turn off high voltage power supplies and de-mate the external high voltage connections at the high voltage supplies.

Warning:  Failure to disconnect the High Voltage prior to pump down could damage test hardware.

High Voltage Power Supplies turned off and disconnected .

Performed by _________ (initials)

Time_________________ 

[image: image2.wmf]Figure 3 -- Engineering "B" Test TV Time Line

Tom Riley

6-0712

11/22/02

TVTL1.skf

, Check Setup in Air

, Bell Closed, PreVac in N2

, No Vac in N2

Close Bell

Vac

Beak Vac

Open Bell

HV Off

HV Off

HV Off, No Transition Tests

5.5 Cycles in about 144 Hours

HV Off

HV On

HV Off

Manned 24/7

11 - "B" Tests

Notes

Survival

Operation

+45 C

+35C

-30 C

-60C

Test B

Test B

Test B

Test B

Test B

Test B

Test B

Test B

Test B

1. HV on, Wait 15 min

2. Data Take, duration 1 hr

3. Release for transition, 10 min

HV Res Test 

6 - High Resistor Tests

Test B

Test B

Test B

6 Hr Min.

HV On

HV On

HV On

HV On

HV On

HV On

HV On

HV On

HV Res Test 

Set Up

3 Cycles

3 Cycles


10.3.1 Pump down Vacuum Chamber

This step will be performed by  Code 545 personnel. 

Initiate vacuum chamber pump down.  The target pressure for this test is 1x10-6 torr.

10.3.2   Reconnect HV Cable

After the chamber pressure has been less than 1 x 10-6 for 4 hours,  
Reconnect the External High Voltage and check that the common return is on the center feed-thru.

Warning:  Miss-connection of the External High Voltage Cable could damage test hardware.

10.4  Thermal Cycling

10.4.1 Following the profile in Figure 3, set the chamber shroud temperature +50 C.

Note: During hot soaks the chamber pressure may go above 1x 10-6 torr due to outgassing.  Therefore it is prerequisite to wait until the chamber pressure is less than 5 x10-5 torr prior to initiating hot soaks.

10.4.2. After all TCs have reached +35 C +/- 2 C, and the chamber pressure is less than 5 x 10-5, begin Hot Soak #1.  Set chamber temperature to +35 C. Record the Hot Soak #1 start time and chamber pressure.

__________Start Time Hot Soak #1
_________________Chamber Pressure (torr)

10.4.3. After 4 hours of Hot Soak # 1, using TDA Test B Procedure Section 3., run Test B.

__________Start Time for Test B

__________End Time for Test B

10.4.4. Hot to Cold Transition  (Following the profile in Figure 3 )

10.4.5. Set the chamber shroud temperature to -45 C.

10.4.6. After all TCs have reached -30 C +/- 2 C, Set chamber shroud temperature to –30 C. Begin Cold Soak #1.

__________Start Time Cold Soak #1

10.4.7. After 2 hours at –30 C +/-2 C, using TDA Test B Procedure Section 3., run Test B.

__________Start Time for Test B

__________End Time for Test B

10.4.8. Cold to Hot Transition  (Following the profile in figure 2 )

10.4.9. Set the chamber shroud temperature to +50 C.

10.4.10. After all TCs have reached +35 C +/- 2 C. Set chamber temperature to +35 C. When the chamber pressure is less than 5 x 10-5, begin Hot Soak #2.  Record the Hot Soak #2 start time and chamber pressure.

__________Start Time Hot Soak #2

_________________Chamber Pressure(torr)

10.4.11. After 2 hours at Hot Soak #2, using TDA Test B Procedure Section 3.0, run Test B.

__________Start Time for Test B

__________End Time for Test B

10.4.12. Hot to Cold Transition  (Following the profile in Figure 3 )

10.4.13. Set the chamber shroud temperature to -45 C.

10.4.14. After all TCs have reached -30 C +/- 2 C, Set chamber temperature to -45 C, begin Cold Soak #2.

__________Start Time Cold Soak #2

10.4.15. After 2 hours at –30 C +/-2 C, using TDA Test B Procedure Section 3.0, run Test B.

__________Start Time for Test B

__________End Time for Test B

10.4.16. Cold to Hot Transition  (Following the profile in Figure 3 )

10.4.17. Set the chamber shroud temperature to +50 C.

10.4.18. After all TCs have reached +35 C +/- 2 C, Set chamber temperature to +35 C. Begin Hot Soak #3.

Record the Hot Soak #3 start time and chamber pressure.

__________Start Time Hot Soak #3
_________________Chamber Pressure(torr)

10.4.19. After 2 hours at Hot Soak #3 C, using TDA Test B Procedure Section 3.0, run Test B.

__________Start Time for Test B

__________End Time for Test B

10.4.20. Hot to Cold Transition  (Following the profile in figure 2 )

10.4.21. Set the chamber shroud temperature to -45 C.

10.4.22. After all TCs have reached -30 C +/- 2 C, Set chamber temperature to -35 C. Begin Cold Soak #3.

__________Start Time Cold Soak #3

10.4.23. After 2 hours at –30 C +/-2 C, using TDA Test B Procedure Section 3.0, run Test B.

__________Start Time for Test B

__________End Time for Test B

10.4.24. Cold Operation to Cold Survival Transition  (Following the profile in Figure 3 )

10.4.25. Verify that the PMT survival heater on the PMT test assembly is turned on.

10.4.26. Set the chamber shroud temperature to -80 C.

10.4.27. After all TCs, (other than those associated with the PMTs, which will be maintained above -40C by survival heaters), have reached -60 C +/- 2 C,  set chamber temperature to -60 C . Begin Cold Survival Soak #1. Record the start time for survival soak #1.  After 4 hours at –60 C +/-2 C, record the end time for Survival Cold Soak #1.

__________Start Time Cold Survival Soak #1 ________End Time for Survival Soak#1

10.4.28. Cold to Hot Transition  (Following the profile in figure 2 )

10.4.29. Set the chamber shroud temperature to +50 C.

10.4.30. After all TCs have reached +45 C +/- 2 C, set chamber temperature to +45 C, begin Hot Survival Soak #1.

__________Start Time Hot Survival Soak #1

10.4.31. After 4 hours at +45 C +/-2 C.  Record the end time for Hot Survival Soak #1.

__________End Time for Hot Survival Soak #1
10.4.32. Hot Survival to Cold Survival Transition  (Following the profile in figure 2 )

10.4.33. Set the chamber shroud temperature to -80 C.

10.4.34. After all TCs, (other than those associated with the PMTs, which will be maintained above       -40C by survival heaters), have reached -60 C +/- 2 C, set chamber temperature to -60 C begin Cold Survival Soak #2. Record the start time for Survival Soak #2.  After 4 hours at –60 C +/-2 C, record the end time for Survival Cold Soak #2.

__________Start Time Cold Survival Soak #2 ________End Time for Survival Soak#2

10.4.35. Cold to Hot Transition  (Following the profile in figure 2 )

10.4.36. Set the chamber shroud temperature to +50  C.

10.4.37. After all TCs have reached +45 C +/- 2 C, Set chamber temperature to +45 C, begin Hot Survival Soak #2.

__________Start Time Hot Survival Soak #2

10.4.38. After 4 hours at +45 C +/-2 C.  Record the end time for Hot Survival Soak #2.

__________End Time for Hot Survival Soak #2
10.4.39. Hot Survival to Cold Operational Transition  (Following the profile in figure 2 )

10.4.40. Set the chamber shroud temperature to -45 C.

10.4.41. After all TCs have reached -30 C +/- 2 C, set chamber temperature to +35 C, begin Cold Soak #4. Record the start time for Cold Soak #4.  After 2 hours at –30 C +/-2 C, Using TDA Test B Procedure Section 4.0, run Test B.

__________Start Time Cold Soak #4 

__________Start time for Test B.

10.4.42. Cold Operation to Cold Survival Transition  (Following the profile in Figure 2 )

10.4.43. Set the chamber shroud temperature to -80 C.

10.4.44. After all TCs, (other than those associated with the PMTs, which will be maintained above       -40C by survival heaters), have reached -60 C +/- 2 C, Set chamber temperature to -60 C, begin Cold Survival Soak #3. Record the start time for survival soak #1.  After 4 hours at –60 C +/-2 C, record the end time for Survival Cold Soak #3.

__________Start Time Cold Survival Soak #3 ________End Time for Cold Survival Soak #3

10.4.45. Cold to Hot Transition  (Following the profile in figure 2 )

10.4.46. Set the chamber shroud temperature to +50 C.

10.4.47. After all TCs have reached +45 C +/- 2 C, set chamber temperature to +45 C, begin Hot Survival Soak #3.

__________Start Time Hot Survival Soak #3

10.4.48. After 4 hours at +45 C +/-2 C.  Record the end time for Hot Survival Soak #3.

__________End Time for Hot Survival Soak #3
10.4.49. Hot Survival to Hot Operational Transition  (Following the profile in Figure 3 )

10.4.50. Set the chamber shroud temperature to +50 C.

10.4.51. After all TCs have reached +35 C +/- 2 C, set chamber temperature to +35 C begin Hot Soak #4. Record the start time for Hot Soak #4.  After 2 hours at +35 C +/-2 C, Using TDA Test B Procedure Section 4.0, run Test B.

__________Start Time Hot Soak #4

__________Start time for Test B.

10.4.52. Set the chamber shroud temperature to +20 C.

10.4.53. After all TCs, have reached 25 C +/- 2 C, begin  Ambient Soak. Record the start time for Ambient Soak.  

10.5        End Thermal Cycling/Break Vacuum
10.5.1
Turn off High Voltage to PMTs and disconnect leads from the power supply.

_____________(initials)

_____________time High Voltage turned off.

10.5.2
After 2 hours at 20 C +/-2 C, backfill the chamber with GN2 for 20 min.  

This step is to be performed by Code 545 personnel.

10.6 Final Functional Test

10.6.1
Turn on the High Voltage to the PMTs and using After 20 minutes of GN2 purge, using TDA Test B Procedure Section 4.0, run Test B.

___________Start time for Test B

10.6.2
Turn of High Voltage to PMTs and disconnect leads from the power supply.

_____________(initials)

_____________time High Voltage turned off.

__________Start Time Cold Survival Soak #3 ________End Time for Cold Survival Soak#3

10.7
Perform High Resistance Test on PMT/Resistor Networks.

Using the PMT/Resistor Network Test Procedure,  check the resistance between the high voltage 

+ input and high voltage return (shield for high voltage connectors), for each PMT/ Resistor Network.  Log resistances below.

Warning:  Failure to connect the High Voltage connections properly could damage test hardware.

PMT

(+)Probe

(-) Probe
Expected

Resistance

_____

Center Conductor
Shield

886 +/-88mohms
_________ Mohms 

_____

Center Conductor
Shield

886 +/-88mohms
_________ Mohms

_____

Center Conductor
Shield

886 +/-88mohms
_________ Mohms

_____

Center Conductor
Shield

886 +/-88mohms
_________ Mohms

10.8
Turn off High Voltage Supplies, Turn off and disconnect all test equipment.

10.8.1. Open the Vacuum Chamber.
Code 545 personnel shall perform this step.

10.8.2. Back Up Data to disks

Appendix A, Cable Drawings

[image: image3.wmf]Engineering "B" Test 

Data Cables

Tom Riley

6-0712

11/12/02

BtestCab1.skf

6 SMA Feed-Thru

5 - TNC to BNC Adapters

5 - BNC to MicroDot

5 - TNC to BNC Adapters

80" RG 178

BNC

Lemo

Not Available

RG 174 Pigtails 

On Units Under Test

T-Vac Chamber

MicroDot

Channel 1

Channel 2

Channel 3

Channel 4

Channel 5

To Data Logger

1

2

3

4

5

1

2

3

4

5

Cable 2

Cable 4

Page 

1

 of 

3

Internal

External


[image: image4.wmf]Engineering "B" Test 

High Voltage Cables

T-Vac Chamber

Cast Block

TP1 HV Rtn

Cable 3

80" RG 59

Safety

Cover

6 Conductor Feed-Thru, 30 KVolt 16 Gauge

6 - 16 Gauge Loose Pins 

      with Heat Shrink

6 - 16 Gauge Loose Pins 

      with Heat Shrink

High Voltage Probe

1. External HV Cable High-Pot

 to 1400 Volts for 20 minutes

HV Rtn on Center Pin

HV Rtn on Center Pin

 HV "T"

Flight HV Cable

Pigtails to

Units Under Test

1

2

3

4

5

Shields

Cable 1

Notes:

To HV Power Supply

Shields

Internal

External

2. Secure internal cable with head srink 

and flight tie wraps.

3. Secure external cable with tye wraps.

Page 

2

 of 

3

Amp 51426-5

Amp 51426-5


[image: image5.wmf]Engineering "B" Test 

Miscellaneous Cables

T-Vac Chamber

MS 3106A-28-21S

37 Female

Safety Gnd

to Non-Flight Rail

t

M1

M2

Metal Body Gnd

Survival Heater

with Thermostat

Ground Straps

Cable 5

60" 18 Gauge Wire

Banana Plug

Red

Black

g

h

s

g

h

s

MS 3106A-28-21P

37 male

60" 18 Gauge Wire

Cable7

Cable 6

37 Conductor Feed-Thru

Page 

3

 of 

3

Internal

External



CHANGE HISTORY LOG

	Revision / Author
	Effective Date
	Description of Changes

	Tom Riley
	10/08/02
	Base Version

	      Bill Reaves
	       10/29/02
	                                         Draft

	      Bill Reaves
	11/15/02
	                                     Final Draft

	Tom Riley
	11/22/02
	Revise illustrations, add cable drawings

	
	
	

	
	
	














TC7





TC6





TTD PMT





TC5











TC4





TC3





TC2





TC5





TC10





TC9





TC8





TC1





PMT


Fixture





TC





TDA 





TTA:13”





Angle is clamped to hard points on the chamber wall and supports the tile.  Tile is 13”x14”.  Angle can be whatever cross-sectional area is available.





8”





16” long





            Test support stand


      -0.25” diameter x 6”  tubes 


       -0.25”x1” cross-sectional                 	area of frame pieces


- 2-21” long sides (other side   is 16”long).  Frame is 3-sided.





Tile rests on bottom of chamber








 Page 1 of 19

