	[image: image1.jpg]



	Document #
	Date Effective

	
	LAT-DS-00739-1
	11/30/01

	
	Prepared by(s)
	Supersedes

	
	Bob Hartman
	None

	GLAST LAT Technical Document
	
	

	
	Subsystem/Office

	
	Anticoincidence Detector Subsystem

	Document Title

	Specifications for ACD Photomultiplier Tubes


Gamma-ray Large Area Space Telescope (GLAST)

Large Area Telescope (LAT)

Specifications for ACD Photomultiplier Tubes

	ORDER FOR SUPPLIES OR SERVICES
	PAGE       OF   PAGES

	IMPORTANT: Mark all packages and papers with contract and/or order numbers.
	1
	10

	1. DATE OF ORDER
	2. CONTRACT NO. (If any)
	6. SHIP TO:

	NOV.  30, 2001
	
	a. NAME OF CONSIGNEE

	3. ORDER NO.
	4. REQUISITION/REFERENCE NO.
	NASA/GODDARD SPACE FLIGHT CENTER     

	S 60463 G
	740-09742A(1P)
	b. STREET ADDRESS

	5. ISSUING OFFICE (Address correspondence to)
	GREENBELT ROAD     

	NASA’s Goddard Space Flight Center

ATTN:  MS. VERONICA C. STUBBS

GREENBELT, MD  20771
	c. CITY
	d. STATE
	e. ZIP CODE

	
	GREENBELT     
	MD
	20771

	
	f. SHIP VIA

	
	     

	
	8. TYPE OF ORDER

	7. TO:
	 FORMCHECKBOX 
 a. PURCHASE
REFERENCE YOUR:           .

Please furnish the following on the terms and conditions specified on both sides of this order and on the attached sheets, if any, including delivery as indicated.
	 FORMCHECKBOX 
 b. DELIVERY - Except for billing instructions on the  reverse, this delivery order is subject to instructions contained on this side only of this form and is issued       subject to the terms and conditions of the above-numbered contract.

	a. NAME OF CONTRACTOR
	
	

	HAMAMATSU CORPORATION


	
	

	b. COMPANY NAME
	
	

	      HAMAMATSU CORPORATION
	
	

	c. STREET ADDRESS
	
	

	360 FOOTHILL ROAD
	
	

	d. CITY
	e. STATE
	f. ZIP CODE
	g. PPC CODE
	h. STATE CODE

	BRIDGEWATER
	NJ
	08807
	GF
	     

	9. ACCOUNTING AND APPROPRIATION DATA
	10. REQUISITIONING OFFICE

	JON: 743-785-20-31-02;APP: 801/20110(01); BLI: A701

OC: 74-2550; AMT:$16,000.00; BN/C: 202
	     

	11. BUSINESS CLASSIFICATION (Check appropriate box(es))

	
 FORMCHECKBOX 
 a. SMALL

 FORMCHECKBOX 
 b. OTHER THAN SMALL

 FORMCHECKBOX 
 c. DISADVANTAGED

 FORMCHECKBOX 
 d. WOMEN-OWNED

	12. F.O.B. POINT
	14. GOVERNMENT B/L NO.
	15. DELIVER TO F.O.B.POINT ON
       OR BEFORE (Date)
	16. DISCOUNT TERMS

	NASA/GODDARD SPACE FLIGHT CENTER
	
	
	

	13. PLACE OF
	     
	MARCH 1, 2002
	

	a. INSPECTION
	b. ACCEPTANCE
	
	
	

	
	
	
	
	

	17. SCHEDULE (See reverse for Rejections)

	ITEM
NO.
(A)
	
SUPPLIES OR SERVICES
(B)
	QUANTITY
ORDERED
©
	
UNIT
(D)
	UNIT
PRICE
(E)
	
AMOUNT
(F)
	QUANTITY
ACCEPTED
(G)

	1
	IN ACCORDANCE WITH THE ATTACHED STATEMENT OF WORK , ADDITIONAL TERMS,  AND THE CONTRACTOR’S PROPOSAL DATED OCTOBER 8, 2001, THE CONTRACTOR SHALL PROVIDE PHOTOMULTIPLIER TUBES (PMT’S) TO BE USED IN THE GLAST ACD
	10
	EACH
	$1,600.00
	$16,000.00
	

	
	SEE ATTACHED FOR ADDITIONAL TERMS
	
	
	
	
	

	
	     
	
	
	     
	     
	

	
	     
	
	     
	     
	     
	

	
	     
	     
	     
	     
	     
	

	SEE BILLING INSTRUCTIONS ON REVERSE
	18. SHIPPING POINT
	19. GROSS SHIPPING WEIGHT
	20. INVOICE NO.
	$16,000.00
	(
	17(h)TOT. (Cont. pages)

	
	     
	     
	     
	
	
	

	
	21. MAIL INVOICE TO:
	
	
	

	
	a. NAME                               
	
	
	

	
	NASA GODDARD SPACE FLIGHT CENTER
	
	
	

	
	b. STREET ADDRESS (or P.O. Box
	$16,000.00
	(
	17(i)

GRAND TOTAL

	
	ACCOUNTS  PAYABLE  CODE 155
	
	
	

	
	c. CITY     
	d. STATE
	e. ZIP CODE


	
	
	

	
	GREENBELT
	MD
	20771
	
	
	

	22. UNITED STATES OF
      AMERICA BY (Signature)
	(
	
	
	23. NAME (Typed)

	
	
	
	
	JAMES S. KING

	
	
	
	DATE
	TITLE: CONTRACTING  OFFICER


 




ORDER NUMBER S-60463-G

ADDITIONAL TERMS

Accounting and Appropriation Data

PCN:  740-09742A(1P)

JON:  743-785-20-31-02

APP:
801/20110(01)



BLI:
A701








OC:
74-2550






AMT:
$16,000.00






TOTAL:  $16,000.00

BNC:
  202

ADDITIONAL TERMS

Delivery Schedule 

1. Basic Contract--Ten (10) Qualification Unit PMT's shall be delivered on or before

 March 1, 2002. 

2. Option 1--Forty (40) Engineering Unit PMT's shall be delivered on or before June 1, 2002. 

3. Option 2--Two hundred ten (210) Flight Unit PMT's shall be delivered within ten (10) months after the exercise of this option.  

Delivery shall be FOB Destination to the Goddard Space Flight Center Greenbelt, MD 20771.
OPTION FOR INCREASED QUANTITY.  THE GOVERNMENT MAY INCREASE THE QUANTITY OF SUPPLIES CALLED FOR IN THE SCHEDULE AT THE UNIT PRICE SPECIFIED IN ACCORDANCE WITH FAR CLAUSE 52.217-6, OPTION FOR INCREASED QUANTITY AS FOLLOWS:


OPTION
DESCRIPTION


QUANTITY
UNIT PRICE
TOTAL PRICE

   1

Engineering Unit PMT
   40

  $1,640.00
 $65,600.00




Tested to Level 3

2 Flight Models Tested to

Level 3

         210 
  $1,350.00
 $283,500.00

Note:  We expect to change the quantity from 210 to 240 to allow sufficient spares. 

NO CHANGES ARE TO BE MADE TO THIS PURCHASE ORDER WITHOUT THE PROPER AUTHORIZATION FROM GSFC PROCUREMENT.

CONTRACTUAL INQUIRIES SHOULD BE DIRECTED TO VERONICA C. STUBBS, CODE 216, GSFC, GREENBELT, MD 20771, PHONE (301) 286-8386 OR FAX (301) 286-1773, EMAIL Veronica.C.Stubbs.1@gsfc.nasa.gov
PRICE PER WRITTEN QUOTE DATED OCTOBER 8, 2001 BY ANDREW ALLEN OF HAMAMATSU CORPORATION

VENDOR'S TAX IDENTIFICATION NUMBER: 13-2638233.

EQUIPMENT USER: DR. ROBERT HARTMAN, CODE 661, PHONE, (301) 286-7178.

THE CONTRACTOR'S INVOICE SHALL CITE ITS TAX IDENTIFICATION NUMBER.

STATEMENT OF WORK/SPECIFICATIONS

GLAST ACD PHOTOMULTIPLIER TUBES

The Gamma-ray Large Area Space Telescope (GLAST) Anticoincidence Detector (ACD) requires 230 small photomultiplier tubes. Photomultiplier tubes are the sensor of choice for the low light levels coming from the plastic scintillators. The high efficiency required for anticoincidence demands a photomultiplier tube with capability of detecting single photoelectrons.  All tubes must be sufficiently reliable and tested for space flight use. 

This procurement is for Photomultiplier Tubes that meet the Photomultiplier Tube (PMT) Requirements detailed below, to be delivered according to the schedule detailed elsewhere in this contract. 

All test data shall be delivered with each PMT, either electronically (preferred) or by paper copy.

Failure reports shall be submitted on any PMT's that fail to meet the Performance Requirements after testing (see below).

PHOTOMULTIPLIER TUBE REQUIREMENTS

PHYSICAL REQUIREMENTS FOR THE PHOTOMULTIPLIER TUBES 

1. Head-on type cylindrical tube 

2. Bi-alkali cathode 

3. Outside diameter less than or equal to 15.2 mm 

4. Length less than or equal to 80 mm, not including leads 

5. Flying lead connections at least 33 mm long, with ends tin-lead solderable to within 10 mm from the glass stem 

6. Minimum cathode diameter 9 mm 

7. As delivered, all PMT's shall be soldered to a temporary base to enable initial electrical testing. 

PERFORMANCE REQUIREMENTS FOR THE PHOTOMULTIPLIER TUBES 

1. Minimum current gain at maximum high voltage shall be 2,000,000 (2E+6). 

2. Quantum efficiency shall be greater than 15.5% (15% for the qualification tubes) at a wavelength of 490 nm. 

3. Anode dark current shall be less than 10 nanoamps at a gain of 5E+5 and a temperature of 20 C. 

4. High voltage required at a gain of 2E+6 shall not exceed 1250 V. 

5. Projected gain degradation shall be less than 30% (84% probability) after 50,000 hours of operation at a mean anode current of 30 nanoamps. The projection shall be based on accelerated life tests at higher currents with tubes of the same design. Proposals shall detail how such tests have been conducted. Except as required below, the PMT's to be delivered under this solicitation shall not undergo such extended accelerated life tests.

ENVIRONMENTAL REQUIREMENTS TO BE MET BY THE PHOTOMULTIPLIER TUBES 

1. Temperature range: The PMT's shall operate within specification over a range ‑30 C to +45 C. 

2. Temperature rate (dT/dt): the PMT's shall operate within specification while experiencing temperature change (within the above range) at a rate of (flight) 5 C/hr; and (test) 20 C/hr. 

3. Temporal stability: At a constant temperature of 20 C, gain variation shall be less than 2% per 10 days. 

4. Vibration: The PMT's shall operate within specification, and with parameter changes not exceeding the values shown below, after experiencing a Delta2 Qualification Level vibration environment in accordance with http://arioch.gsfc.nasa.gov/302/gevs-se/sec2-4.pdf. 

5. Shock: The PMT's shall operate within specification, and with parameter changes not exceeding the values shown below, after experiencing Delta2 Qualification Level shock testing in accordance with http://arioch.gsfc.nasa.gov/302/gevs-se/sec2-4.pdf. 

6. Radiation: the units shall operate within specification, and with parameter changes not exceeding the values shown below, after a total dose of 10 kRad. 

Limits on Performance Changes Due to Vibration, Shock, or Radiation 

Quantum efficiency shall not change by more than 10%; 

Current gain shall not change by more than 25%; 

Anode dark current (at a temperature of 20 C) shall not increase by more than 100% if it is initially greater than or equal to 1 nA.  If the initial anode dark current is less than 1 nA, it shall not increase by more than a factor of 10 after vibration, shock, or radiation.

(All three of the parameters listed above must be within specification both before and after the vibration, shock, or radiation exposure.) 

QUALITY REQUIREMENTS TO BE MET BY THE PHOTOMULTIPLIER TUBES 

1. The Photomultiplier Tube Quality Plan shall be consistent with the following: 

ISO-9001; 

GSFC 433-MAR-0001, "The Mission Assurance Requirements for the Large Area Telescope;" 

LAT Performance Assurance Implementation Plan (SLAC LAT-MD-00039); 

"Quality Plan for the Anticoincidence Detector" (GSFC ACD-QA-8001). 

2. The PMT's must be space flight qualifiable according to 433-SPEC-0001 (GLAST, Mission System Specification). 

TESTING - QUALIFICATION UNIT PMT'S (LEVEL 2 TESTING):

1) Visual Inspection

2) DC Measurements (quantum efficiency at 490 nm, anode luminous sensitivity at 1250V, derived current gain at 1250V, nominal voltage for gain of 5E+5, anode dark current at gain of 5E+5). Any PMT that fails to meet the performance requirements listed above shall be rejected.  Failure reports are not required on PMT's rejected at this point.

3) Random vibration - The ten (10) Qualification Unit PMT's shall be vibration tested by the contractor to the Delta2 Qualification Level specified in http://arioch.gsfc.nasa.gov/302/gevs-se/sec2-4.pdf. The PMT's shall meet the above performance requirements both before and after the vibration test. Furthermore, changes in performance parameters following the vibration test shall not exceed the limits shown above under "Limits on Performance Changes Due to Vibration, Shock, or Radiation." 

4) Visual Inspection (window or envelope defects, loose particles, internal structure).

5) DC Measurements as in 2) - Any PMT that fails to meet the performance requirements listed above shall be rejected.  Failure reports are required on PMT's rejected at this point.

6) 12-hour Burn-in

7) Visual Inspection (repeat item #1)

8) DC Measurements as in 2) - Any PMT that fails to meet the performance requirements listed above shall be rejected.  Failure reports are required on PMT's rejected at this point.

All test data (from steps 2, 5, and 8 above) shall be delivered at the same time a PMT is delivered.

The Qualification Unit PMT's shall not be subjected to the 100-hour burn-in described below. 

TESTING – ENGINEERING AND FLIGHT UNIT PMT'S (LEVEL 3 TESTING):

1) Visual Inspection (window or envelope defects, loose particles, internal structure)

2) DC Measurements - (quantum efficiency at 490 nm, anode luminous sensitivity at 1250V, derived current gain at 1250V, nominal voltage for gain of 5E+5, anode dark current at gain of 5E+5). Any PMT that fails to meet the performance requirements listed above shall be rejected.  Failure reports are not required on PMT's rejected at this point.

3) Random vibration -

Following the testing described above, each PMT shall be vibration tested at Delta2 Flight Acceptance Level in accordance with http://arioch.gsfc.nasa.gov/302/gevs-se/sec2-4.pdf 

4) Visual Inspection (repeat item #1)

5) DC Measurements (quantum efficiency at 490 nm, anode luminous sensitivity at 1250V, derived current gain at 1250V, nominal voltage for gain of 5E+5, anode dark current at gain of 5E+5)  

Failure reports shall be submitted on any Engineering Unit or Flight Unit PMT's that fail to meet the Performance Requirements, or which demonstrate performance changes in excess of the limits shown above under "Limits on Performance Changes Due to Vibration, Shock, or Radiation" after vibration testing. Such failure shall be reported to the Technical Officer within two (2) working days. The failed PMT shall undergo failure analysis, the results of which shall be delivered at a mutually agreed upon time, based on the type of failure. 

6) 100-hour Burn-in - Following the testing described above, each Flight Unit PMT shall be operated at 30 microamp anode current for 100 hours (gain drift characterization), with the Current Amplification measured initially and after 12, 24, 48 and 100 hours of operation.

7) Visual Inspection (repeat item #1)

8) DC Measurements - The full set of performance parameters (quantum efficiency at 490 nm, anode luminous sensitivity at 1250V, derived current gain, nominal voltage for gain of 5E+5, anode dark current at gain of 5E+5).  Any PMT shall be rejected, and a Failure Report submitted, if its post-burn-in Cathode Luminous Sensitivity, Anode Dark Current (at 20 C temperature), or Current Gain differs from the pre-test value by more the limits shown above under "Limits on Performance Changes Due to Vibration, Shock, or Radiation", or if it no longer meets the Performance Requirements shown above.  Furthermore, any PMT whose gain has decreased by more than 10% after the first 12 hours, or by more than 25% after the full 100 hours, shall be rejected; a Failure Report shall be submitted on such a PMT.

All test data from steps 2, 5, and 8, as well as the gain data from the gain tests at 0, 12, 24, 48, and 100 hours, shall be delivered at the time each Flight Unit PMT is delivered.

PAGE  
7

